Bio-Medical Equipment Service Co.

M1400 Learn Code Instructions

(  CMS Central  Revision D and above

- Press the key labeled INSTRUMENT CONFIG.

- Select TELEMETRY FREQUENCY key.

- Enter the following pass code 14432.

- The screen being displayed for you now is a list of all of the channels for that central.  Each channel represents a sector on your screen.

- Locate the sector you are working with by scrolling up and down with your keypad until you have the channel you are working with highlighted.

- Press CONFIRM.

- Press CONFIRM one more time.

- Put a known good battery in your transmitter and connect transmitter to a patient simulator.

- Now press the key labeled LEARN CODE.

- Now press the nurse call button on your transmitter.

- You should now have a signal on that sector.

- Hook transmitter up to a patient simulator and admit the patient for that sector.

(   CMS Central Station Revision D and Below

- Press key labeled PATIENT WINDOW.

- You will now see arrows pop up beside each sector.  Press the corresponding bed key for the sector you need to learn the code for.

- Next press key labeled MORE CONTROLS.

- Press the key labeled MORE CONTROLS one more time.

- Put a known good battery in your transmitter and connect transmitter to a patient simulator.

- Select key labeled LEARN CODE.

- Now press the nurse call button on your transmitter.

- You should now have a signal on that sector.

- Hook transmitter up to a patient simulator and admit the patient for that sector.